# Software Requirements Specification

## Accountability Log Application (ALA)

### Medical Treatment Tracking System

| Document Information |  |
| --- | --- |
| Document ID | ALA-SRS-001 |
| Version | 4.0 |
| Date | January 2026 |
| Status | Draft |
| Classification | Internal |
| Safety Classification | IEC 62304 Class B |
| FDA Documentation Level | Enhanced |

## Approvals

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Author |  |  |  |
| Technical Review |  |  |  |
| QA Review |  |  |  |
| Project Manager |  |  |  |

## Revision History

| Version | Date | Author | Description |
| --- | --- | --- | --- |
| 1.0 | 2024 | - | Initial SRS for ALA |
| 2.0 | December 2025 | - | Complete rewrite reflecting current implementation |
| 3.0 | December 2025 | - | Full overhaul: Added risk integration (ISO 14971), HIPAA 2025 compliance, cybersecurity section, missing features (attachments, treatment types, radioactivity), populated traceability matrix |
| 4.0 | January 2026 | - | Major update: Added Offline Mode/PWA (Section 3.14), Treatment Continuation (Section 3.15), Enhanced Removal Procedure with Discrepancy Tracking (Section 3.8), new hazards (HAZ-011 to HAZ-016), updated data dictionary |

## Table of Contents

- Introduction

- Overall Description

- System Features
- 3.1-3.13: Core Features
- 3.14: Offline Mode & PWA Architecture
- 3.15: Treatment Continuation

- External Interface Requirements

- Non-Functional Requirements

- Data Requirements

- Verification and Validation

- Risk Management Integration

- Cybersecurity Requirements

- Appendices

## 1. Introduction

### 1.1 Purpose

This Software Requirements Specification (SRS) document describes the functional and non-functional requirements for the Accountability Log Application (ALA), version 2.0. This document is intended for:

- Software developers implementing and maintaining the system

- Quality Assurance personnel validating system compliance

- Project managers overseeing development

- Clinical operations personnel using the system

- Regulatory affairs for compliance documentation

### 1.2 Scope

The Accountability Log Application (ALA) is a web-based medical treatment tracking system designed to replace paper-based accountability logs (QSR-4001-01) for seed applicator procedures used in brachytherapy treatments.

In Scope:

- User authentication and authorization via Priority ERP integration

- Treatment session management (insertion and removal procedures)

- Real-time applicator tracking with 8-state workflow management

- Barcode/QR code scanning for applicator identification

- Digital signature capture for treatment finalization

- PDF generation for regulatory documentation

- Comprehensive audit logging for regulatory compliance

- Integration with Alpha Tau Medical's Priority ERP system

Out of Scope:

- Medical device control or operation

- Patient medical records management

- Billing or insurance processing

- Inventory management (handled by Priority ERP)

### 1.3 Product Overview

The ALA system digitizes the tracking of seed applicator usage during brachytherapy procedures. The system consists of:

- Frontend: React/TypeScript single-page application optimized for mobile and tablet use

- Backend: Express/TypeScript REST API server

- Database: PostgreSQL for treatment and audit data persistence

- Integration: Priority ERP OData API for patient data, site access, and applicator validation

### 1.4 Definitions, Acronyms, and Abbreviations

| Term | Definition |
| --- | --- |
| ALA | Accountability Log Application |
| Applicator | Medical device used to deploy radioactive seeds during brachytherapy |
| Brachytherapy | Cancer treatment using radioactive seeds implanted near tumors |
| ERP | Enterprise Resource Planning |
| HttpOnly Cookie | Secure cookie type inaccessible to JavaScript (XSS protection) |
| JWT | JSON Web Token - standard for secure authentication tokens |
| OData | Open Data Protocol - REST-based data access protocol |
| Position Code | Priority ERP user classification (99 = admin) |
| Priority | Alpha Tau Medical's ERP system |
| QR Code | Quick Response Code - 2D barcode format |
| Seed | Radioactive implant used in brachytherapy |
| Subject ID | Unique patient treatment identifier from Priority ORDERS |
| Treatment | A single brachytherapy procedure (insertion or removal) |

### 1.5 References

| ID | Document | Description |
| --- | --- | --- |
| REF-01 | IEEE/ISO/IEC 29148:2018 | Systems and software engineering - Life cycle processes - Requirements engineering |
| REF-02 | IEC 62304:2006+AMD1:2015 | Medical device software - Software life cycle processes |
| REF-03 | ISO 14971:2019 | Medical devices - Application of risk management to medical devices |
| REF-04 | QSR-4001-01 | Accountability Log (paper form being replaced) |
| REF-05 | Priority ERP OData API | Integration specification for Priority system |
| REF-06 | OWASP Top 10 | Web application security risks |
| REF-07 | HIPAA Security Rule 2025 | Updated cybersecurity requirements (Federal Register 2025-01-06) |
| REF-08 | FDA Cybersecurity Guidance | Content of Premarket Submissions for Management of Cybersecurity in Medical Devices |

### 1.6 Document Conventions

- Requirements are identified using format: SRS-[MODULE]-[NUMBER]

- Priority levels: Critical, High, Medium, Low

- "Shall" indicates mandatory requirements

- "Should" indicates recommended but optional requirements

- "May" indicates optional features

## 2. Overall Description

### 2.1 Product Perspective

The ALA operates as a standalone web application that integrates with Alpha Tau Medical's existing Priority ERP system. It does not replace Priority but extends it with specialized treatment tracking capabilities.

System Context:

┌─────────────────┐     ┌─────────────────┐     ┌─────────────────┐

│   Mobile/Tablet │────▶│   ALA Frontend  │────▶│   ALA Backend   │

│   (Browser)     │     │   (React SPA)   │     │   (Express API) │

└─────────────────┘     └─────────────────┘     └────────┬────────┘

│

┌────────────────────────────────┼────────────────┐

│                                │                │

▼                                ▼                ▼

┌───────────────┐              ┌─────────────────┐ ┌─────────────┐

│  PostgreSQL   │              │  Priority ERP   │ │ Email SMTP  │

│   Database    │              │   OData API     │ │   Service   │

└───────────────┘              └─────────────────┘ └─────────────┘

### 2.2 Product Functions

The ALA provides the following high-level functions:

| Function | Description |
| --- | --- |
| Authentication | Email-based verification with JWT session management |
| Treatment Selection | Choose patient treatments from Priority ERP data |
| Applicator Scanning | QR code/barcode scanning with validation |
| 8-State Workflow | Track applicator status through procedure stages |
| Progress Tracking | Real-time visualization of treatment completion |
| Treatment Finalization | Digital signature capture and PDF generation |
| Removal Tracking | Track seed removal procedures (30+ days post-insertion) |
| Audit Logging | Comprehensive audit trail for regulatory compliance |

### 2.3 User Classes and Characteristics

| User Class | Role | Permissions | Technical Expertise |
| --- | --- | --- | --- |
| Hospital Worker | Clinical staff performing treatments | View/edit treatments at assigned sites only | Basic - familiar with medical procedures |
| AlphaTau Employee | Clinical operations support | View/edit all treatments, approve applicator exceptions | Intermediate - trained on ALA system |
| Administrator | System management | Full access including logs and configuration | Advanced - IT/technical background |

Position Code 99: Users with Priority position code 99 receive full administrative access regardless of assigned role.

### 2.4 Operating Environment

| Component | Requirement |
| --- | --- |
| Client Devices | Modern web browser (Chrome 90+, Safari 14+, Firefox 88+, Edge 90+) on mobile phones, tablets, or desktop computers |
| Screen Size | Minimum 320px width (responsive design) |
| Network | Internet connection required for Priority ERP integration |
| Camera | Device camera required for QR code scanning (optional manual entry available) |
| Server | Azure Virtual Machine (Ubuntu 20.04 LTS) |
| Database | PostgreSQL 14+ |
| Runtime | Node.js 18 LTS |

### 2.5 Design and Implementation Constraints

| Constraint | Description |
| --- | --- |
| C-01 | System shall integrate with existing Priority ERP OData API without modifications to Priority |
| C-02 | All patient-identifiable data shall remain in Priority ERP; ALA stores only Subject IDs |
| C-03 | System shall be accessible on mobile devices in clinical environments |
| C-04 | System shall support concurrent users performing treatments at different sites |
| C-05 | System shall maintain audit trail for regulatory compliance (IEC 62304) |
| C-06 | Authentication tokens shall be stored in HttpOnly secure cookies (XSS protection) |

### 2.6 Assumptions and Dependencies

Assumptions:

- Users have valid Priority ERP accounts with appropriate site access

- Clinical environment has WiFi or cellular data connectivity

- Applicators have valid QR codes/barcodes matching Priority inventory

- Users are trained on brachytherapy procedures and ALA usage

Dependencies:

- Priority ERP system availability for authentication and data

- Azure cloud infrastructure availability

- Email delivery service for verification codes

- PostgreSQL database availability

### 2.7 Regulatory Context (IEC 62304)

Software Safety Classification: Class B

Rationale: The ALA is a documentation and tracking system that does not directly control medical device operation. However, incorrect tracking could lead to:

- Unreported applicator usage

- Incorrect seed counts in patient documentation

- Compliance documentation errors

The system implements the following safety measures:

- Comprehensive audit logging of all status changes

- Validation against Priority ERP master data

- Digital signature requirements for finalization

- Immutable audit trail for regulatory review

## 3. System Features

### 3.1 Authentication and Authorization

#### 3.1.1 Description and Priority

Priority: Critical

The system shall authenticate users via email verification codes and authorize access based on Priority ERP user data and role assignments.

#### 3.1.2 Stimulus/Response Sequences

Login Flow:

- User enters email address

- System validates email against Priority PHONEBOOK

- System sends 6-digit verification code via email

- User enters verification code

- System validates code and creates authenticated session

- User redirected to procedure selection

Session Management:

- Authenticated user makes request

- System validates JWT from HttpOnly cookie

- If valid, request proceeds

- If expired/invalid, user redirected to login

#### 3.1.3 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-AUTH-001 | The system shall validate user email addresses against Priority ERP PHONEBOOK table | Critical |
| SRS-AUTH-002 | The system shall generate cryptographically secure 6-digit verification codes | Critical |
| SRS-AUTH-003 | The system shall send verification codes via email using configured SMTP service | Critical |
| SRS-AUTH-004 | Verification codes shall expire after 10 minutes | High |
| SRS-AUTH-005 | The system shall hash verification codes using bcrypt before storage | Critical |
| SRS-AUTH-006 | The system shall issue JWT tokens with 7-day expiration upon successful verification | High |
| SRS-AUTH-007 | JWT tokens shall be stored in HttpOnly secure cookies (not accessible to JavaScript) | Critical |
| SRS-AUTH-008 | The system shall track failed verification attempts per user | High |
| SRS-AUTH-009 | The system shall allow users to request code resend | Medium |
| SRS-AUTH-010 | The system shall extract user site access from Priority metadata | Critical |
| SRS-AUTH-011 | Users with Position Code 99 shall receive full access to all sites | High |
| SRS-AUTH-012 | The system shall implement automatic session timeout after inactivity | Medium |
| SRS-AUTH-013 | The system shall provide logout functionality that clears authentication cookies | High |
| SRS-AUTH-014 | The system shall prevent user enumeration by returning generic error messages | High |

### 3.2 Treatment Selection

#### 3.2.1 Description and Priority

Priority: Critical

The system shall allow users to select treatment procedures (insertion or removal) from Priority ERP data based on their site access permissions.

#### 3.2.2 Stimulus/Response Sequences

Insertion Selection:

- User selects "Insertion" procedure type

- System fetches available treatments from Priority ORDERS

- System filters treatments by user's authorized sites

- User selects patient/treatment from list

- System creates local treatment record linked to Priority

Removal Selection:

- User selects "Removal" procedure type

- System fetches completed insertion treatments eligible for removal

- User selects treatment to perform removal

- System creates removal treatment linked to original insertion

#### 3.2.3 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-TSEL-001 | The system shall allow users to choose between insertion and removal procedure types | Critical |
| SRS-TSEL-002 | The system shall fetch treatment data from Priority ERP ORDERS table | Critical |
| SRS-TSEL-003 | The system shall display only treatments at sites the user is authorized to access | Critical |
| SRS-TSEL-004 | The system shall combine related Priority procedures for pancreas treatments | High |
| SRS-TSEL-005 | The system shall allow treatment search by Subject ID, date, and site | High |
| SRS-TSEL-006 | The system shall display patient identifier (DETAILS field) from Priority | High |
| SRS-TSEL-007 | For removal procedures, the system shall only show insertions completed at least 30 days prior | High |
| SRS-TSEL-008 | The system shall create a local treatment record with reference to Priority ID | Critical |
| SRS-TSEL-009 | The system shall store treatment metadata: type, site, date, Subject ID, surgeon | High |
| SRS-TSEL-010 | The system shall record the user who created the treatment | High |

### 3.3 Applicator Scanning and Validation

#### 3.3.1 Description and Priority

Priority: Critical

The system shall scan and validate applicator QR codes/barcodes against Priority ERP inventory data.

#### 3.3.2 Stimulus/Response Sequences

Scan Flow:

- User initiates barcode scan (camera or manual entry)

- System decodes QR code to extract serial number

- System validates serial number against Priority SIBD_APPLICATUSELIST

- System checks for duplicate scanning

- System displays applicator details and validation result

- User confirms or rejects applicator

#### 3.3.3 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-SCAN-001 | The system shall support QR code scanning using device camera | Critical |
| SRS-SCAN-002 | The system shall support manual serial number entry as fallback | Critical |
| SRS-SCAN-003 | The system shall validate applicator serial numbers against Priority SIBD_APPLICATUSELIST | Critical |
| SRS-SCAN-004 | The system shall detect and notify if an applicator was already scanned for the current treatment | High |
| SRS-SCAN-005 | The system shall detect applicators assigned to different treatments | High |
| SRS-SCAN-006 | The system shall display applicator details: serial number, seed quantity, applicator type | High |
| SRS-SCAN-007 | The system shall retrieve seed quantity from Priority based on serial number | Critical |
| SRS-SCAN-008 | The system shall allow users to manually override seed quantity with justification | Medium |
| SRS-SCAN-009 | The system shall record package label from scanned data when available | Medium |
| SRS-SCAN-010 | The system shall store applicator type (PARTS.PARTDES) from Priority | Medium |

### 3.4 Applicator 8-State Workflow Management

#### 3.4.1 Description and Priority

Priority: Critical

The system shall track applicators through an 8-state workflow that varies by treatment type.

#### 3.4.2 Stimulus/Response Sequences

State Transition (Pancreas/Prostate):

- Applicator scanned → Status: SEALED

- User opens package → Status: OPENED

- User loads applicator → Status: LOADED

- User deploys seeds → Status: INSERTED (or FAULTY/DISCHARGED/DEPLOYMENT_FAILURE)

State Transition (Skin):

- Applicator scanned → Status: SEALED

- User deploys seeds → Status: INSERTED (or FAULTY)

#### 3.4.3 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-WKFL-001 | The system shall support 8 applicator statuses: SEALED, OPENED, LOADED, INSERTED, FAULTY, DISPOSED, DISCHARGED, DEPLOYMENT_FAILURE | Critical |
| SRS-WKFL-002 | The system shall enforce treatment-specific transition rules (Pancreas/Prostate vs. Skin vs. Generic) | Critical |
| SRS-WKFL-003 | For Pancreas/Prostate: SEALED→OPENED→LOADED→(INSERTED¦DISCHARGED¦DEPLOYMENT_FAILURE) | Critical |
| SRS-WKFL-004 | For Skin: SEALED→(INSERTED¦FAULTY) | Critical |
| SRS-WKFL-005 | Terminal statuses (INSERTED, FAULTY, DISPOSED, DISCHARGED, DEPLOYMENT_FAILURE) shall not allow further transitions | Critical |
| SRS-WKFL-006 | The system shall require comments for all failure terminal statuses (FAULTY, DISPOSED, DISCHARGED, DEPLOYMENT_FAILURE) | High |
| SRS-WKFL-007 | The system shall create audit log entries for every status transition | Critical |
| SRS-WKFL-008 | Audit logs shall record: old status, new status, user, timestamp, reason, request ID | Critical |
| SRS-WKFL-009 | The system shall display available transitions based on current status and treatment type | High |
| SRS-WKFL-010 | The system shall visually distinguish status types: success (green), failure (dark), in-progress (yellow/red) | Medium |

### 3.5 Real-Time Progress Tracking

#### 3.5.1 Description and Priority

Priority: High

The system shall display real-time progress of treatment completion.

#### 3.5.2 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-PROG-001 | The system shall display total applicators vs. processed applicators | High |
| SRS-PROG-002 | The system shall display total seeds vs. inserted seeds | High |
| SRS-PROG-003 | The system shall calculate and display completion percentage | High |
| SRS-PROG-004 | The system shall show usage type distribution (full, faulty, none) | Medium |
| SRS-PROG-005 | The system shall show applicator status breakdown by state | Medium |
| SRS-PROG-006 | Progress display shall update in real-time as applicators are processed | High |

### 3.6 Treatment Finalization and Digital Signatures

#### 3.6.1 Description and Priority

Priority: Critical

The system shall support treatment finalization with digital signature verification.

#### 3.6.2 Stimulus/Response Sequences

Hospital Auto-Signature:

- Hospital user clicks "Finalize"

- System auto-captures user information as signature

- PDF generated with signature

- Treatment marked complete

AlphaTau Verification:

- User initiates finalization

- User selects signer from site users list

- System sends verification code to signer's email

- Signer enters verification code

- System validates code (3 attempts max, 1-hour expiration)

- PDF generated with verified signature

- Treatment marked complete

#### 3.6.3 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-FINL-001 | The system shall support two finalization types: hospital_auto and alphatau_verified | Critical |
| SRS-FINL-002 | Hospital users may auto-sign treatments using their authenticated identity | High |
| SRS-FINL-003 | AlphaTau finalization shall require verification code sent to signer's email | Critical |
| SRS-FINL-004 | Signature verification codes shall expire after 1 hour | High |
| SRS-FINL-005 | The system shall limit signature verification attempts to 3 per code | High |
| SRS-FINL-006 | The system shall fetch available signers from Priority PHONEBOOK for the treatment site | High |
| SRS-FINL-007 | The system shall record signer name, email, position, and signature timestamp | Critical |
| SRS-FINL-008 | Finalization shall mark treatment as complete with completedBy and completedAt | Critical |
| SRS-FINL-009 | Completed treatments shall be immutable (no further applicator changes) | Critical |

### 3.7 PDF Generation and Export

#### 3.7.1 Description and Priority

Priority: High

The system shall generate comprehensive treatment documentation in PDF format.

#### 3.7.2 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-EXPRT-001 | The system shall generate PDF documents containing complete treatment information | High |
| SRS-EXPRT-002 | PDF shall include: treatment summary, site, date, Subject ID, surgeon | High |
| SRS-EXPRT-003 | PDF shall include complete applicator list with serial numbers, statuses, seed quantities | High |
| SRS-EXPRT-004 | PDF shall include signature information (name, position, timestamp) | Critical |
| SRS-EXPRT-005 | PDF shall include seed count totals (inserted, removed, remaining) | High |
| SRS-EXPRT-006 | The system shall store generated PDF in database (BYTEA field) | High |
| SRS-EXPRT-007 | The system shall support JSON export of treatment data | Medium |
| SRS-EXPRT-008 | Exported documents shall not include internal Alpha Tau employee names | Medium |

### 3.8 Removal Workflow

#### 3.8.1 Description and Priority

Priority: High

The system shall support tracking of seed removal procedures with comprehensive discrepancy tracking to document any difference between inserted and removed sources.

#### 3.8.2 Stimulus/Response Sequences

Removal Procedure Flow:

- User selects removal procedure for eligible treatment (30+ days post-insertion)

- System loads applicators from original insertion

- User marks applicators as removed (group removal)

- User optionally tracks individual seed removals

- System auto-calculates total sources removed

- If discrepancy detected (removed ≠ inserted), user clarifies discrepancy

- User completes removal procedure form

- User finalizes treatment

Discrepancy Clarification Flow:

- System detects removed sources ≠ inserted sources

- System displays discrepancy clarification section

- User selects applicable categories (Lost, Retrieved to site, Remained in tissue, Other)

- User enters amounts per category

- System validates sum of amounts equals sources not removed

- User provides comments per category as needed

#### 3.8.3 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-RMVL-001 | The system shall allow removal procedures for treatments at least 30 days post-insertion | High |
| SRS-RMVL-002 | The system shall display applicators from the original insertion treatment | High |
| SRS-RMVL-003 | The system shall track number of seeds removed per applicator | High |
| SRS-RMVL-004 | The system shall allow removal comments per applicator | Medium |
| SRS-RMVL-005 | The system shall track removal timestamp | High |
| SRS-RMVL-006 | The system shall track who performed each removal | High |
| SRS-RMVL-007 | The system shall display removal progress (seeds removed vs. total seeds) | High |
| SRS-RMVL-008 | The system shall color-code final count: green if zero seeds remain, red otherwise | Medium |
| SRS-RMVL-009 | The system shall provide a removal procedure form for capturing removal details | High |
| SRS-RMVL-010 | Removal procedure form shall capture: removal date, all-same-date flag | High |
| SRS-RMVL-011 | If sources removed on different dates, system shall capture additional date and reason | High |
| SRS-RMVL-012 | The system shall auto-calculate total sources removed from applicators + individual seeds | High |
| SRS-RMVL-013 | The system shall detect discrepancy when removed sources ≠ inserted sources | Critical |
| SRS-RMVL-014 | Discrepancy clarification shall track 4 categories: Lost, Retrieved to site, Remained in tissue (removal failure), Other | High |
| SRS-RMVL-015 | Each discrepancy category shall include: checked flag, amount, and comment fields | High |
| SRS-RMVL-016 | Sum of clarified amounts must equal the number of sources not removed | High |
| SRS-RMVL-017 | The system shall support individual seed removal with predefined reasons | Medium |
| SRS-RMVL-018 | Individual seed removal reasons shall include: Applicator rupture, Patient arrived from home, Other | Medium |
| SRS-RMVL-019 | The system shall calculate removal progress from applicator removals + individual seed removals | High |
| SRS-RMVL-020 | The system shall allow treatment completion with missing sources when properly documented | Medium |
| SRS-RMVL-021 | The system shall store discrepancy clarification data as JSON | High |
| SRS-RMVL-022 | The system shall store individual seed removal notes as JSON | Medium |

### 3.9 Audit Logging

#### 3.9.1 Description and Priority

Priority: Critical

The system shall maintain comprehensive audit logs for regulatory compliance.

#### 3.9.2 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-AUDT-001 | The system shall log all applicator status transitions in ApplicatorAuditLog table | Critical |
| SRS-AUDT-002 | Audit log entries shall include: applicator ID, old status, new status, user, timestamp | Critical |
| SRS-AUDT-003 | Audit log entries shall include reason/comments when provided | High |
| SRS-AUDT-004 | Audit log entries shall include request ID for correlation | High |
| SRS-AUDT-005 | Audit logs shall be immutable (no updates or deletes) | Critical |
| SRS-AUDT-006 | The system shall index audit logs by applicator ID and timestamp for efficient retrieval | High |
| SRS-AUDT-007 | The system shall support audit log queries by time range and applicator | High |

### 3.10 Admin Dashboard

#### 3.10.1 Description and Priority

Priority: Medium

The system shall provide administrative dashboard for system monitoring.

#### 3.10.2 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-ADMN-001 | The system shall provide dashboard accessible only to admin users | Medium |
| SRS-ADMN-002 | Dashboard shall display treatment statistics | Medium |
| SRS-ADMN-003 | Dashboard shall display system logs | Medium |
| SRS-ADMN-004 | Dashboard shall support filtering by date range and site | Low |

### 3.11 File Attachments

#### 3.11.1 Description and Priority

Priority: Medium

The system shall support file attachments per applicator for documenting images, certificates, or other relevant files.

#### 3.11.2 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-ATCH-001 | The system shall support file attachments per applicator | Medium |
| SRS-ATCH-002 | The system shall track attachment filename for each upload | Medium |
| SRS-ATCH-003 | The system shall track attachment count per applicator | Medium |
| SRS-ATCH-004 | The system shall track attachment size in bytes | Medium |
| SRS-ATCH-005 | The system shall track attachment sync status (pending, syncing, synced, failed) | Medium |

### 3.12 Treatment Type Handling

#### 3.12.1 Description and Priority

Priority: Critical

The system shall support multiple treatment types with type-specific workflow behaviors.

#### 3.12.2 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-TYPE-001 | The system shall support treatment types: insertion, removal, pancreas_insertion, prostate_insertion, skin_insertion | Critical |
| SRS-TYPE-002 | Pancreas treatments shall combine related Priority orders via reference chain detection | High |
| SRS-TYPE-003 | Treatment priorityId shall support JSON array format for combined/referenced orders | High |
| SRS-TYPE-004 | Each treatment type shall have type-specific workflow transitions per SRS-WKFL-002 | Critical |

### 3.13 Radioactivity Tracking

#### 3.13.1 Description and Priority

Priority: High

The system shall track radioactivity levels (activity per seed) for treatment documentation.

#### 3.13.2 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-RAD-001 | The system shall track activity per seed (radioactivity level) for treatments | High |
| SRS-RAD-002 | Activity per seed shall be included in PDF export per SRS-EXPRT-002 | High |
| SRS-RAD-003 | Activity per seed shall be retrieved from Priority ORDERS data | High |

### 3.14 Offline Mode & PWA Architecture

#### 3.14.1 Description and Priority

Priority: Critical

The system shall implement Progressive Web App (PWA) capabilities enabling offline operation with HIPAA-compliant data encryption and automatic synchronization when connectivity is restored.

#### 3.14.2 Stimulus/Response Sequences

Offline Download Flow:

- User selects treatment for offline use

- System validates user has site access

- System downloads treatment bundle (treatment + applicators)

- System encrypts PHI using AES-256-GCM

- System stores encrypted data in IndexedDB

- Bundle marked with 24-hour expiry

Offline Operation Flow:

- Network connectivity lost (or user goes offline)

- System displays offline status banner

- User can view downloaded treatments

- User can add/update applicators offline

- Changes queued with SHA-256 integrity hashes

- Finalization BLOCKED until online

Sync Flow (Reconnection):

- Network connectivity restored

- System auto-triggers sync

- Pending changes pushed to server (idempotent)

- Server validates and detects conflicts

- Conflicts stored for resolution

- Medical-critical conflicts require admin resolution

#### 3.14.3 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-OFFL-001 | The system shall implement Progressive Web App (PWA) architecture with service workers | Critical |
| SRS-OFFL-002 | The system shall use IndexedDB for offline data storage via Dexie.js | Critical |
| SRS-OFFL-003 | The system shall encrypt PHI at rest using AES-256-GCM (HIPAA 2025 compliant) | Critical |
| SRS-OFFL-004 | The system shall use PBKDF2 key derivation with 100,000 iterations (OWASP 2024) | Critical |
| SRS-OFFL-005 | The system shall detect and track network connectivity status in real-time | High |
| SRS-OFFL-006 | The system shall queue offline changes with SHA-256 integrity hashes | Critical |
| SRS-OFFL-007 | The system shall auto-sync pending changes when network connectivity restores | High |
| SRS-OFFL-008 | The system shall implement exponential backoff retry (1s initial, 60s max, 5 retries) | High |
| SRS-OFFL-009 | The system shall support treatment bundle download for offline use | High |
| SRS-OFFL-010 | Offline bundles shall expire after 24 hours | High |
| SRS-OFFL-011 | The system shall detect and store sync conflicts for resolution | Critical |
| SRS-OFFL-012 | Medical-critical status conflicts (INSERTED, FAULTY) shall require admin resolution | Critical |
| SRS-OFFL-013 | The system shall maintain offline audit log with device ID and timestamps | Critical |
| SRS-OFFL-014 | Treatment finalization shall be BLOCKED while offline (requires digital signature) | Critical |
| SRS-OFFL-015 | The system shall support clock synchronization with server (5-minute tolerance) | High |
| SRS-OFFL-016 | The system shall track and display storage usage statistics | Medium |
| SRS-OFFL-017 | The system shall warn users at 80% storage capacity | Medium |
| SRS-OFFL-018 | The system shall auto-delete expired offline bundles | Medium |
| SRS-OFFL-019 | The system shall request persistent storage permission from browser | Medium |
| SRS-OFFL-020 | The system shall display offline status banner with pending change count | High |
| SRS-OFFL-021 | The system shall provide manual sync trigger button | Medium |
| SRS-OFFL-022 | The system shall implement user-controlled app updates (no auto-update during treatment) | Critical |
| SRS-OFFL-023 | Encrypted fields shall include: patientName, subjectId, surgeon, serialNumber, comments, removalComments | Critical |
| SRS-OFFL-024 | The system shall support idempotent sync using change hashes (prevent duplicate processing) | High |
| SRS-OFFL-025 | The system shall warn iOS users about 7-day storage eviction risk | Medium |

### 3.15 Treatment Continuation

#### 3.15.1 Description and Priority

Priority: High

The system shall allow users to continue a completed insertion treatment within a 24-hour window, enabling additional applicators to be added in follow-up sessions.

#### 3.15.2 Stimulus/Response Sequences

Continuation Detection Flow:

- User selects patient for insertion procedure

- System checks for completed treatments with same patient + site

- System evaluates 24-hour eligibility window (from lastActivityAt or completedAt)

- If eligible treatment found, system displays continuation modal

- Modal shows: hours remaining, reusable applicator count, parent details

- User confirms continuation or cancels

Continuation Creation Flow:

- User confirms continuation

- System creates new treatment with parentTreatmentId linking to parent

- New treatment inherits: type, subjectId, patientName, site, surgeon, priority info

- New treatment uses current date (not parent date)

- User navigates to treatment documentation to add applicators

- Continuation treatment finalized separately from parent

#### 3.15.3 Functional Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-CONT-001 | The system shall allow continuation of completed insertion treatments within 24-hour window | High |
| SRS-CONT-002 | Continuation eligibility shall be based on lastActivityAt or completedAt timestamp | High |
| SRS-CONT-003 | The system shall track parent-child treatment relationship via parentTreatmentId | High |
| SRS-CONT-004 | Continuation treatment shall inherit: type, subjectId, patientName, site, surgeon, priority info | High |
| SRS-CONT-005 | Continuation treatment shall use current date (not parent treatment date) | High |
| SRS-CONT-006 | The system shall display modal when continuable treatment detected | Medium |
| SRS-CONT-007 | Continuation modal shall show: hours remaining, reusable applicator count, parent details | Medium |
| SRS-CONT-008 | Reusable applicators are those with OPENED or LOADED status | High |
| SRS-CONT-009 | Only insertion treatments can be continued (removal treatments cannot be continued) | High |
| SRS-CONT-010 | The system shall provide API endpoints for continuation eligibility check and creation | High |

## 4. External Interface Requirements

### 4.1 User Interfaces

#### 4.1.1 General UI Requirements

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-UI-001 | The system shall be responsive and function on mobile devices (320px minimum width) | High |
| SRS-UI-002 | The system shall use accessible color contrast ratios (WCAG 2.1 AA) | Medium |
| SRS-UI-003 | The system shall provide clear error messages for all user actions | High |
| SRS-UI-004 | The system shall display loading indicators during async operations | Medium |
| SRS-UI-005 | The system shall display environment indicator (staging/production) | Medium |

#### 4.1.2 Key Screens

| Screen | Purpose | Key Elements |
| --- | --- | --- |
| Login | User authentication | Email input, verification code input |
| Procedure Type | Select insertion/removal | Two prominent selection buttons |
| Treatment Selection | Choose patient/treatment | Search filters, treatment list, selection |
| Treatment Documentation | Scan and process applicators | Scanner, applicator form, progress display |
| Use List | Review applicators, finalize | Applicator table, edit, finalize button |
| Seed Removal | Track removal procedure | Applicator list, removal inputs, progress |

### 4.2 Hardware Interfaces

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-HW-001 | The system shall access device camera for QR code scanning via browser APIs | High |
| SRS-HW-002 | The system shall function without camera using manual entry | Critical |

### 4.3 Software Interfaces

#### 4.3.1 Priority ERP OData API

| Endpoint | Purpose | Data |
| --- | --- | --- |
| PHONEBOOK | User authentication, site access | Email, position code, authorized sites |
| ORDERS | Treatment/patient data | Subject ID, patient details, site, date |
| SIBD_APPLICATUSELIST | Applicator inventory | Serial number, seed quantity, applicator type |

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-API-001 | The system shall authenticate to Priority API using configured credentials | Critical |
| SRS-API-002 | The system shall handle Priority API timeouts (30 second default) | High |
| SRS-API-003 | The system shall cache Priority data to reduce API calls | Medium |
| SRS-API-004 | The system shall provide fallback test data when Priority is unavailable (development only) | Medium |

#### 4.3.2 Database Interface

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-DB-001 | The system shall use PostgreSQL 14+ for data persistence | Critical |
| SRS-DB-002 | The system shall use Sequelize ORM for database operations | High |
| SRS-DB-003 | The system shall implement database connection pooling | High |
| SRS-DB-004 | The system shall implement health checks for database connectivity | High |

### 4.4 Communication Interfaces

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-COM-001 | The system shall use HTTPS for all client-server communication | Critical |
| SRS-COM-002 | The system shall use SMTP for sending verification emails | High |
| SRS-COM-003 | The system shall implement email retry logic for delivery failures | Medium |

## 5. Non-Functional Requirements

### 5.1 Performance Requirements

| ID | Requirement | Priority | Metric |
| --- | --- | --- | --- |
| SRS-PERF-001 | Page load time shall be under 3 seconds on 4G connection | High | < 3000ms |
| SRS-PERF-002 | API response time shall be under 2 seconds for standard operations | High | < 2000ms |
| SRS-PERF-003 | QR code scanning shall decode within 500ms | High | < 500ms |
| SRS-PERF-004 | PDF generation shall complete within 10 seconds | Medium | < 10000ms |
| SRS-PERF-005 | System shall support 50 concurrent users | Medium | 50 users |
| SRS-PERF-006 | Offline bundle download shall complete within 30 seconds | High | < 30000ms |
| SRS-PERF-007 | Offline sync operation shall complete within 60 seconds | High | < 60000ms |
| SRS-PERF-008 | IndexedDB read/write operations shall complete within 500ms | High | < 500ms |
| SRS-PERF-009 | PHI encryption/decryption shall complete within 100ms per record | High | < 100ms |

### 5.2 Safety Requirements (IEC 62304)

| ID | Requirement | Priority | Rationale |
| --- | --- | --- | --- |
| SRS-SAFE-001 | The system shall validate all applicator data against Priority master data | Critical | Prevent incorrect tracking |
| SRS-SAFE-002 | The system shall require confirmation before terminal status transitions | High | Prevent accidental state changes |
| SRS-SAFE-003 | The system shall maintain immutable audit trail of all status changes | Critical | Regulatory compliance |
| SRS-SAFE-004 | The system shall prevent modification of finalized treatments | Critical | Data integrity |
| SRS-SAFE-005 | The system shall require comments for failure statuses | High | Documentation of issues |
| SRS-SAFE-006 | The system shall display clear warnings for unusual conditions | High | User awareness |

### 5.3 Security Requirements

| ID | Requirement | Priority | Implementation |
| --- | --- | --- | --- |
| SRS-SEC-001 | The system shall store JWT tokens in HttpOnly secure cookies | Critical | XSS protection |
| SRS-SEC-002 | The system shall hash verification codes using bcrypt | Critical | Password security |
| SRS-SEC-003 | The system shall implement rate limiting on authentication endpoints | Critical | Brute-force protection |
| SRS-SEC-004 | The system shall implement rate limiting on API endpoints | High | DoS protection |
| SRS-SEC-005 | The system shall enforce HTTPS in production | Critical | Transport security |
| SRS-SEC-006 | The system shall implement CORS origin validation | High | Cross-origin protection |
| SRS-SEC-007 | The system shall implement security headers via Helmet.js | High | Security hardening |
| SRS-SEC-008 | The system shall prevent user enumeration via generic error messages | High | Information disclosure |
| SRS-SEC-009 | The system shall validate all user input | Critical | Injection prevention |
| SRS-SEC-010 | The system shall use AES-256 encryption for data at rest | Critical | Data protection |
| SRS-SEC-011 | The system shall use TLS 1.3 for all data in transit | Critical | Transport security |
| SRS-SEC-012 | Multi-factor authentication shall be used for ePHI access (email verification code) | Critical | HIPAA 2025 |
| SRS-SEC-013 | Network segmentation shall isolate ePHI systems | High | HIPAA 2025 |
| SRS-SEC-014 | Audit logs shall be retained for minimum 6 years per HIPAA | Critical | Regulatory compliance |
| SRS-SEC-015 | Audit log integrity shall be verified via immutability controls | High | Data integrity |
| SRS-SEC-016 | Session timeout shall occur after 30 minutes of inactivity | High | Access control |

Rate Limiting Configuration:

| Endpoint Type | Limit | Window |
| --- | --- | --- |
| Authentication | 20 requests | 15 minutes |
| Code Request | 10 requests | 15 minutes |
| Verification | 15 requests | 15 minutes |
| Token Validation | 50 requests | 15 minutes |
| General API | 100 requests | 15 minutes |

### 5.4 Software Quality Attributes

| Attribute | Requirement |
| --- | --- |
| Availability | System shall be available 99.5% during business hours |
| Reliability | System shall recover from crashes within 5 minutes |
| Maintainability | System shall use TypeScript for type safety |
| Portability | System shall run in Docker containers |
| Scalability | System shall support horizontal scaling via Docker Swarm |

### 5.5 Business Rules

| ID | Rule | Priority |
| --- | --- | --- |
| SRS-BR-001 | Removal procedures are only allowed 30+ days after insertion | High |
| SRS-BR-002 | Position Code 99 users have full site access | High |
| SRS-BR-003 | Test data (test@example.com) is isolated from production | Critical |
| SRS-BR-004 | Failure statuses require explanatory comments | High |
| SRS-BR-005 | Completed treatments cannot be modified | Critical |

### 5.6 Backup and Recovery Requirements

| ID | Requirement | Priority | Metric |
| --- | --- | --- | --- |
| SRS-RECV-001 | The system shall support automated database backup every 24 hours | High | Daily backup |
| SRS-RECV-002 | Recovery Time Objective (RTO) shall be less than 4 hours | High | < 4 hours |
| SRS-RECV-003 | Recovery Point Objective (RPO) shall be less than 24 hours | High | < 24 hours |
| SRS-RECV-004 | Backup integrity shall be verified weekly | Medium | Weekly verification |

## 6. Data Requirements

### 6.1 Logical Data Model

┌─────────────┐       ┌─────────────────┐       ┌──────────────┐

│    Users    │       │   Treatments    │       │  Applicators │

├─────────────┤       ├─────────────────┤       ├──────────────┤

│ id (PK)     │──┐    │ id (PK)         │──┐    │ id (PK)      │

│ name        │  │    │ type            │  │    │ serialNumber │

│ email       │  │    │ subjectId       │  │    │ seedQuantity │

│ phoneNumber │  │    │ patientName     │  │    │ usageType    │

│ role        │  └───▶│ userId (FK)     │  └───▶│ treatmentId  │

│ metadata    │       │ completedBy (FK)│       │ status       │

└─────────────┘       │ site            │       │ addedBy (FK) │

│ date            │       │ comments     │

│ isComplete      │       └──────┬───────┘

└─────────────────┘              │

│

┌────────────────────┐     ┌─────────────────────┐    │

│  TreatmentPdfs     │     │ ApplicatorAuditLog  │    │

├────────────────────┤     ├─────────────────────┤    │

│ id (PK)            │     │ id (PK)             │◀───┘

│ treatmentId (FK)   │     │ applicatorId (FK)   │

│ pdfData (BYTEA)    │     │ oldStatus           │

│ signatureType      │     │ newStatus           │

│ signerName         │     │ changedBy           │

│ signerEmail        │     │ changedAt           │

└────────────────────┘     │ reason              │

└─────────────────────┘

### 6.2 Data Dictionary

#### 6.2.1 Users Table

| Field | Type | Constraints | Description |
| --- | --- | --- | --- |
| id | UUID | PK, auto-generated | Unique user identifier |
| name | VARCHAR(255) | NOT NULL | User's display name |
| email | VARCHAR(255) | UNIQUE | User's email address |
| phoneNumber | VARCHAR(50) | UNIQUE | User's phone number |
| role | ENUM | NOT NULL, default 'hospital' | User role (hospital, alphatau, admin) |
| verificationCode | VARCHAR(255) | nullable | Hashed verification code |
| verificationExpires | TIMESTAMP | nullable | Code expiration time |
| failedAttempts | INTEGER | NOT NULL, default 0 | Failed verification count |
| lastLogin | TIMESTAMP | nullable | Last successful login |
| metadata | JSON | nullable | Priority user data (position, sites) |

#### 6.2.2 Treatments Table

| Field | Type | Constraints | Description |
| --- | --- | --- | --- |
| id | UUID | PK, auto-generated | Unique treatment identifier |
| type | ENUM | NOT NULL | Treatment type (insertion, removal) |
| subjectId | VARCHAR(100) | NOT NULL, indexed | Priority Subject ID |
| patientName | VARCHAR(255) | nullable | Patient identifier from Priority |
| site | VARCHAR(100) | NOT NULL | Treatment site |
| date | TIMESTAMP | NOT NULL, indexed | Treatment date |
| isComplete | BOOLEAN | NOT NULL, default false | Completion status |
| priorityId | VARCHAR(100) | nullable | Reference to Priority system |
| userId | UUID | FK→users, NOT NULL | Creating user |
| completedBy | UUID | FK→users, nullable | Completing user |
| completedAt | TIMESTAMP | nullable | Completion timestamp |
| seedQuantity | INTEGER | nullable | Total seed count |
| activityPerSeed | FLOAT | nullable | Radiation activity per seed (Bq) |
| surgeon | VARCHAR(255) | nullable | Performing surgeon |
| parentTreatmentId | UUID | FK→treatments, nullable | Links continuation to parent treatment |
| lastActivityAt | TIMESTAMP | nullable | Last applicator activity for 24-hour continuation window |
| removalDate | DATE | nullable | Date of removal procedure |
| allSourcesSameDate | BOOLEAN | nullable | Whether all sources removed on same day |
| additionalRemovalDate | DATE | nullable | Secondary removal date if different |
| reasonNotSameDate | TEXT | nullable | Reason for different removal dates |
| discrepancyClarification | JSON | nullable | Discrepancy tracking data (Lost, Retrieved, Remained, Other) |
| individualSeedsRemoved | INTEGER | nullable, default 0 | Count of individually removed seeds |
| individualSeedNotes | JSON | nullable | Notes for individual seed removals |
| removalGeneralComments | TEXT | nullable | General removal comments |
| syncStatus | ENUM | nullable | Offline sync status (pending, syncing, synced, failed) |
| version | INTEGER | NOT NULL, default 1 | Optimistic locking version for sync conflicts |

Notes:

- The priorityId field may contain a JSON array when treatments combine multiple Priority orders (e.g., pancreas multi-stage treatments).

- The parentTreatmentId links continuation treatments to their parent (24-hour continuation window).

- The discrepancyClarification stores JSON with structure: { lost: {checked, amount, comment}, retrievedToSite: {...}, remainedInTissue: {...}, other: {..., description} }.

#### 6.2.3 Applicators Table

| Field | Type | Constraints | Description |
| --- | --- | --- | --- |
| id | UUID | PK, auto-generated | Unique applicator identifier |
| serialNumber | VARCHAR(100) | NOT NULL, indexed | Applicator serial number |
| seedQuantity | INTEGER | NOT NULL, default 0 | Seeds in applicator |
| usageType | ENUM | NOT NULL, default 'full' | Usage outcome (full, faulty, none) |
| status | VARCHAR(50) | nullable, validated | Workflow status |
| packageLabel | VARCHAR(10) | nullable | Package identifier |
| insertionTime | TIMESTAMP | NOT NULL | When applicator was used |
| comments | TEXT | nullable | Usage comments |
| treatmentId | UUID | FK→treatments, NOT NULL | Parent treatment |
| addedBy | UUID | FK→users, NOT NULL | User who added |
| isRemoved | BOOLEAN | NOT NULL, default false | Removal tracking |
| removalTime | TIMESTAMP | nullable | When removed |
| removedBy | UUID | FK→users, nullable | User who removed |
| applicatorType | VARCHAR(255) | nullable | Type from Priority PARTS |
| attachmentFilename | VARCHAR(255) | nullable | Uploaded file name |
| attachmentFileCount | INTEGER | NOT NULL, default 0 | Number of attached files |
| attachmentSizeBytes | BIGINT | NOT NULL, default 0 | Total attachment size |
| attachmentSyncStatus | ENUM | NOT NULL, default 'pending' | Sync state (pending, syncing, synced, failed) |

#### 6.2.4 ApplicatorAuditLog Table

| Field | Type | Constraints | Description |
| --- | --- | --- | --- |
| id | UUID | PK, auto-generated | Unique log entry identifier |
| applicatorId | UUID | FK→applicators, NOT NULL | Related applicator |
| oldStatus | VARCHAR(50) | nullable | Previous status |
| newStatus | VARCHAR(50) | NOT NULL, validated | New status |
| changedBy | VARCHAR(255) | NOT NULL | User email who made change |
| changedAt | TIMESTAMP | NOT NULL, default NOW() | When change occurred |
| reason | TEXT | nullable | Reason for change |
| requestId | VARCHAR(100) | nullable | Correlation ID |

#### 6.2.5 SyncConflict Table (NEW - Offline Support)

| Field | Type | Constraints | Description |
| --- | --- | --- | --- |
| id | UUID | PK, auto-generated | Unique conflict identifier |
| entityType | ENUM | NOT NULL | Type: 'treatment' or 'applicator' |
| entityId | UUID | NOT NULL, indexed | ID of the conflicting entity |
| localData | JSON | NOT NULL | Data from offline device |
| serverData | JSON | NOT NULL | Data currently on server |
| conflictType | VARCHAR(50) | NOT NULL | Type: 'version_mismatch', 'status_conflict', 'field_conflict' |
| resolution | ENUM | nullable | Resolution: 'local_wins', 'server_wins', 'merged', null (unresolved) |
| resolvedBy | UUID | FK→users, nullable | User who resolved conflict |
| resolvedAt | TIMESTAMP | nullable | When conflict was resolved |
| requiresAdmin | BOOLEAN | NOT NULL, default false | Medical-critical conflicts require admin resolution |
| deviceId | VARCHAR(100) | NOT NULL | Device that created offline change |
| overwrittenData | JSON | nullable | Data that was overwritten (HIPAA audit trail) |
| createdAt | TIMESTAMP | NOT NULL, default NOW() | When conflict was detected |

#### 6.2.6 OfflineAuditLog Table (NEW - Offline Support)

| Field | Type | Constraints | Description |
| --- | --- | --- | --- |
| id | UUID | PK, auto-generated | Unique audit entry identifier |
| entityType | ENUM | NOT NULL | Type: 'treatment' or 'applicator' |
| entityId | UUID | NOT NULL, indexed | ID of the affected entity |
| operation | ENUM | NOT NULL | Operation: 'create', 'update', 'status_change', 'sync' |
| deviceId | VARCHAR(100) | NOT NULL | Device that performed operation |
| userId | UUID | FK→users, NOT NULL | User who performed operation |
| beforeState | JSON | nullable | Entity state before change |
| afterState | JSON | NOT NULL | Entity state after change |
| offlineAt | TIMESTAMP | NOT NULL | When operation occurred offline |
| syncedAt | TIMESTAMP | nullable | When operation was synced to server |
| changeHash | VARCHAR(64) | NOT NULL, indexed | SHA-256 hash for idempotency checking |
| metadata | JSON | nullable | Additional operation metadata |
| createdAt | TIMESTAMP | NOT NULL, default NOW() | When audit entry was created |

### 6.3 Data Integrity and Retention

| ID | Requirement | Priority |
| --- | --- | --- |
| SRS-DATA-001 | The system shall use foreign key constraints for referential integrity | Critical |
| SRS-DATA-002 | The system shall use database transactions for multi-step operations | Critical |
| SRS-DATA-003 | Audit log records shall never be deleted | Critical |
| SRS-DATA-004 | Treatment data shall be retained for minimum 7 years (regulatory requirement) | Critical |
| SRS-DATA-005 | The system shall implement database backups | High |

## 7. Verification and Validation

### 7.1 Verification Methods

| Requirement Category | Verification Method |
| --- | --- |
| Authentication (SRS-AUTH-*) | Unit tests, integration tests, security penetration testing |
| Treatment Selection (SRS-TSEL-*) | Integration tests with Priority API mocks |
| Applicator Scanning (SRS-SCAN-*) | Unit tests, E2E tests with test barcodes |
| Workflow (SRS-WKFL-*) | Unit tests for state machine, integration tests |
| Progress (SRS-PROG-*) | Frontend component tests |
| Finalization (SRS-FINL-*) | E2E tests, manual QA verification |
| Export (SRS-EXPRT-*) | Integration tests, manual PDF review |
| Removal (SRS-RMVL-*) | Integration tests, E2E tests for discrepancy tracking |
| Audit (SRS-AUDT-*) | Database queries, integration tests |
| Security (SRS-SEC-*) | Security scanning, penetration testing |
| Performance (SRS-PERF-*) | Load testing, performance monitoring |
| Offline Mode (SRS-OFFL-*) | PWA testing, offline simulation, sync conflict tests, encryption validation |
| Treatment Continuation (SRS-CONT-*) | Integration tests, 24-hour window validation, E2E tests |

### 7.2 Traceability Matrix

The complete requirements traceability matrix is maintained in a separate document:

See: srs/traceability-matrix.md

The traceability matrix includes:

- All 215 requirements with unique IDs

- Hazard linkage (from ISO 14971 analysis)

- Test case references

- Implementation status

Summary Statistics:

- Total Requirements: 215

- Implemented: 183 (85.1%)

- Needs Verification: 27 (12.6%)

- Pending Implementation: 5 (2.3%)

## 8. Risk Management Integration

This section documents the integration of ISO 14971:2019 risk management with software requirements.

### 8.1 Risk Management File Reference

The complete hazard analysis is maintained in a separate document:

See: srs/hazard-analysis.md

### 8.2 Software Safety Classification

IEC 62304 Safety Class: B

Rationale: The ALA is a documentation and tracking system that does not directly control medical device operation. However, incorrect tracking could result in:

- Unreported applicator usage affecting regulatory compliance

- Incorrect seed counts in patient documentation

- Treatment documentation errors

### 8.3 Hazard Summary

| Hazard ID | Description | Initial Risk | Residual Risk | Related Requirements |
| --- | --- | --- | --- | --- |
| HAZ-001 | Incorrect applicator tracking | Medium | ALARP | SRS-SCAN-003, SRS-AUDT-001 |
| HAZ-002 | Incorrect seed count | Medium | ALARP | SRS-SCAN-007, SRS-SCAN-008 |
| HAZ-003 | Unauthorized access to patient data | Medium | Low | SRS-AUTH-*, SRS-SEC-* |
| HAZ-004 | Data loss during documentation | Medium | ALARP | SRS-DATA-002, SRS-RECV-* |
| HAZ-005 | Audit trail tampering | Medium | ALARP | SRS-AUDT-005, SRS-SEC-015 |
| HAZ-006 | Applicator used on wrong patient | Medium | ALARP | SRS-SCAN-004, SRS-SCAN-005 |
| HAZ-007 | Treatment finalized without verification | Medium | Low | SRS-FINL-003, SRS-FINL-009 |
| HAZ-008 | Removal on wrong treatment | Medium | ALARP | SRS-RMVL-001, SRS-TSEL-007 |
| HAZ-009 | Session hijacking | Medium | Low | SRS-AUTH-007, SRS-SEC-016 |
| HAZ-010 | Brute force attack | Low | Low | SRS-SEC-003, SRS-AUTH-004 |
| HAZ-011 | Data loss during offline sync | Medium | ALARP | SRS-OFFL-006, SRS-OFFL-011 |
| HAZ-012 | Stale offline data used in treatment | Medium | Low | SRS-OFFL-010, SRS-OFFL-015 |
| HAZ-013 | Unauthorized offline PHI access | Medium | ALARP | SRS-OFFL-003, SRS-OFFL-004, SRS-OFFL-023 |
| HAZ-014 | Clock skew causing timestamp errors | Low | Low | SRS-OFFL-015, SRS-AUDT-003 |
| HAZ-015 | Treatment finalized offline without verification | Medium | ALARP | SRS-OFFL-014 |
| HAZ-016 | Continuation on wrong patient | Medium | ALARP | SRS-CONT-003, SRS-CONT-004, SRS-CONT-009 |

### 8.4 Safety-Critical Requirements

The following requirements are classified as safety-critical based on their role in risk control:

| Requirement ID | Description | Related Hazards |
| --- | --- | --- |
| SRS-SCAN-003 | Validate applicator against Priority | HAZ-001, HAZ-006 |
| SRS-SCAN-005 | Detect wrong treatment applicator | HAZ-006 |
| SRS-AUDT-005 | Immutable audit logs | HAZ-001, HAZ-005 |
| SRS-AUTH-007 | HttpOnly secure cookies | HAZ-003, HAZ-009 |
| SRS-FINL-009 | Immutable finalized treatments | HAZ-007 |
| SRS-SEC-003 | Rate limiting on authentication | HAZ-003, HAZ-010 |
| SRS-OFFL-003 | Encrypt PHI at rest using AES-256-GCM | HAZ-013 |
| SRS-OFFL-006 | Queue offline changes with integrity hashes | HAZ-011 |
| SRS-OFFL-011 | Detect and store sync conflicts | HAZ-011 |
| SRS-OFFL-012 | Medical-critical conflicts require admin resolution | HAZ-011, HAZ-012 |
| SRS-OFFL-014 | Block treatment finalization offline | HAZ-015 |
| SRS-CONT-003 | Track parent-child treatment relationship | HAZ-016 |

### 8.5 Risk Control Verification

All risk control measures shall be verified through:

- Code review confirming implementation

- Unit tests for control logic

- Integration tests for end-to-end validation

- Security testing for access controls

## 9. Cybersecurity Requirements

This section addresses FDA cybersecurity guidance and HIPAA 2025 security requirements.

### 9.1 Cybersecurity Framework

| ID | Requirement | Priority | Reference |
| --- | --- | --- | --- |
| SRS-CYBER-001 | The system shall maintain a Software Bill of Materials (SBOM) | High | FDA Guidance |
| SRS-CYBER-002 | The system shall undergo vulnerability scanning every 6 months | High | HIPAA 2025 |
| SRS-CYBER-003 | The system shall undergo annual penetration testing | High | HIPAA 2025 |
| SRS-CYBER-004 | The system shall have documented incident response procedures | High | HIPAA 2025 |
| SRS-CYBER-005 | The system shall have patch management process (critical patches within 30 days) | High | FDA Guidance |

### 9.2 Software Bill of Materials (SBOM)

The system shall maintain an SBOM documenting:

- All open-source dependencies

- Version numbers for each component

- Known vulnerabilities (CVEs)

- License information

### 9.3 Vulnerability Management

| Activity | Frequency | Responsible |
| --- | --- | --- |
| Vulnerability scanning | Every 6 months | DevOps |
| Penetration testing | Annual | Third-party security firm |
| Dependency updates | Monthly review | Development team |
| Critical patch deployment | Within 30 days | DevOps |

### 9.4 Incident Response

The system shall support incident response through:

- Comprehensive audit logging (per SRS-AUDT-*)

- User session management (per SRS-AUTH-*)

- Network monitoring integration capability

- Forensic data preservation

## Appendices

### Appendix A: Applicator Status State Diagram

See docs/srs/applicator-state-diagram.md

┌─────────┐

│ SEALED  │

└────┬────┘

│

┌───────────────┼───────────────┐

│               │               │

▼               ▼               ▼

┌─────────┐     ┌─────────┐     ┌─────────┐

│ OPENED  │     │INSERTED │     │ FAULTY  │

│(Panc/Pr)│     │ (Skin)  │     │ (Skin)  │

└────┬────┘     └─────────┘     └─────────┘

│

┌────┴────┬───────────┐

│         │           │

▼         ▼           ▼

┌────────┐ ┌───────┐ ┌──────────┐

│ LOADED │ │FAULTY │ │ DISPOSED │

└────┬───┘ └───────┘ └──────────┘

│

┌────┴────┬────────────────┬─────────────────┐

│         │                │                 │

▼         ▼                ▼                 ▼

┌────────┐ ┌───────────┐ ┌────────────────────┐

│INSERTED│ │DISCHARGED │ │DEPLOYMENT_FAILURE  │

└────────┘ └───────────┘ └────────────────────┘

Legend:

- Terminal states: INSERTED, FAULTY, DISPOSED, DISCHARGED, DEPLOYMENT_FAILURE

- In-progress states: SEALED, OPENED, LOADED

- Green highlight: INSERTED (success)

- Black highlight: All failure terminal states

### Appendix B: User Role Permission Matrix

| Permission | Hospital | AlphaTau | Admin |
| --- | --- | --- | --- |
| Login | Yes | Yes | Yes |
| View treatments (own sites) | Yes | Yes | Yes |
| View treatments (all sites) | No | Yes | Yes |
| Create insertion treatment | Yes | Yes | Yes |
| Create removal treatment | Yes | Yes | Yes |
| Add applicators | Yes | Yes | Yes |
| Edit applicators | Yes | Yes | Yes |
| Finalize treatment (auto) | Yes | No | Yes |
| Finalize treatment (verified) | N/A | Yes | Yes |
| View admin dashboard | No | No | Yes |
| View system logs | No | No | Yes |

Note: Position Code 99 overrides all restrictions and grants full admin access.

### Appendix C: Priority API Integration Details

#### Authentication

- Method: Basic Authentication

- Credentials: Configured via environment variables

#### PHONEBOOK Query

GET /PHONEBOOK?$filter=EMAIL eq '{email}'

&$select=EMAIL,NAME,POSITION,PHONE,SITES

#### ORDERS Query

GET /ORDERS?$filter=SITE eq '{site}' and DATE ge {startDate}

&$select=ORDNAME,DETAILS,SIESSION,SITE,DATE,SURGEON

#### SIBD_APPLICATUSELIST Query

GET /SIBD_APPLICATUSELIST?$filter=SERIALNUMBER eq '{serialNumber}'

&$select=SERIALNUMBER,SEEDQTY,PARTDES

### Appendix D: Glossary

| Term | Definition |
| --- | --- |
| Applicator | Medical device used to deploy brachytherapy seeds |
| Brachytherapy | Internal radiation therapy using implanted radioactive sources |
| HttpOnly Cookie | Cookie with HttpOnly flag preventing JavaScript access |
| JWT | JSON Web Token for stateless authentication |
| OData | Open Data Protocol for REST APIs |
| Position Code | Priority ERP user classification number |
| Subject ID | Unique patient/treatment identifier in Priority ERP |
| Terminal Status | Applicator status that does not allow further transitions |

## Export to Word Document

To export this document to Word format for QA review:

# Run the conversion script from project root

node docs/convert-to-docx.js

This generates docs/ALA_SRS.docx using the docx npm library.

*Document generated in compliance with IEEE/ISO/IEC 29148:2018 and IEC 62304:2006+AMD1:2015*